2023 NANS Abstract Submission

https://neuromodulation.org/Meetings/2023NANSAnnualMeeting/AbstractFAQ.aspx

Thursday, July 28, 2022, at 1:00 pm (CDT)

Working title: Mechanisms of Increased Charge Injection Capacity of a porous Injectable Electrode known as the Injectrode

Ricardo Sanchez1, James Trevethan1, Kyle Golobish2, Jack Wragan2, Derrick Liu2, Manfred Franke2, Stephan Nieuwoudt2, Andrew Shoffstall2,3, Kip Ludwig1,2

[1] Dept. of Biomedical Engineering, University of Wisconsin, Madison, Wisconsin

[2] Neuronoff, Inc., Cleveland, Ohio, USA

[3] Dept. of Biomedical Engineering, Case Western Reserve University, Cleveland, Ohio

Introduction

Applications of neuromodulation therapies are often limited by the invasiveness of the surgical procedure needed for implantation, high failure rates of implanted devices, and inconsistent therapeutic efficacy due to either off-target effects or lack of therapy at safe charge injection levels. To address several of these limitations we have developed the Injectrode, an injectable electrode designed to minimize surgical invasiveness and device failure while providing a high surface area for safe charge injection. In this work, we seek to understand the mechanisms and limitations of increased charge injection through a high surface area Injectrode formed with a microstructure of thin Pt-Ir wire braided into a flexible helical macroelectrode.

Methods

Electrical Impedance Spectroscopy (EIS), Cyclic Voltammetry (CV), and Voltage Transient (VT) analysis were performed on these electrodes (with a theoretical exposed surface area of 235mm^2) as well as geometric surface area (GSA) and electrochemical surface area (ESA) equivalent electrodes of the same material. Data was analyzed to assess the impact of the high electrochemical surface provided by the fine wire microstructure of the electrode on electrode impedance as well as charge injection capacity (CIC) measured during slow CVs and fast Voltage Transients.

Results

EIS analysis demonstrate the Injectrode has similar faradaic impedance and capacitance to an ESA electrode suggesting the increased surface area of the Injectrode is available for charge Injection. In contrast, slow scan CVs show the Injectrode has a CIC in between that of GSA and ESA electrodes. This suggests a time dependent mechanism for charge injection that we hypothesize may result from diffusion limitations caused by the Injectrode’s braided geometry limiting the electrochemical reactions during the CV analysis. Preliminary results also show tissue ingrowth into the Injectrode microstructure.

Conclusion

The Injectrode  could enable neuromodulation therapy through a minimally invasive procedure with an electrode that conforms to the patients neuroanatomy. This is achieved through a braided fine wire structure, which our electrochemical analysis shows may provide a higher CIC than a GSA equivalent. This increase in charge injection through a braided microstructure may provide an alternative to electrode coatings, which are currently used to increase CIC of neuromodulation electrodes. While the increased CIC with electrode coatings often deteriorates in the presence of biofouling, we hypothesis that a microstructure allowing for tissue growth into the electrode may be more resilient than traditional coating materials to in vivo decreases in CIC caused by biofouling.

References (if any)

Disclosure of Conflicts of Interest

NANS Abstract Requirements:

Title — The title will be published in meeting materials. As an example, a properly formatted title would read: "Programming Optimization Strategies for Burst May Improve Outcomes," not "Programming optimization strategies for burst may improve outcomes."

Abstract Body — This includes an Introduction, Methods, Results, and Conclusion.

Length — Should not exceed 400 words.

References — Include any references that were used in preparing the abstract.

Figures — Charts, graphs, or images can be submitted. File(s) must be in .jpg, .gif, or .png format.

Additional Information — You will need to provide three keywords that describe your abstract.

Disclosure of Conflict of Interest — The Abstract Submitter, Presenting Author, and all Co-Authors of the submitted abstract are required to complete the Disclosure of Conflict of Interest Form, as well as an agreement with the Terms and Conditions.

CME Requirements — Abstracts accepted for the 2023 NANS Annual Meeting cannot contain advertising, logos, trade names, or a product-group message. To ensure all scientific content is free from commercial bias, please include generic names or terms in your abstract content. If the abstract includes trade names, names from several companies should be used to present all available treatment options.

Review Work — All abstract details, including author information, will be published exactly as submitted, without exception. Please review submissions carefully for typographical errors.